## Functional Requirements
EAST BAY MUNICIPAL UTILITY DISTRICT
LIMS RFI Business Requirements |  |  |  | Vendor Response
Requirement ID | Requirements | Priority | Category | Solution Type | Availability | Response
General System Requirements
1. Comply with regulatory requirements and industry standards, not limited to: |  |  | General
1.1. | Good Automated Laboratory Practices (GALP) | Must have
1.2. | International Organization for Standardization: ISO/IEC 17025 | Must have
1.3. | NELAC Institute (TNI) | Must have
· National Environmental Laboratory Accreditation Program (NELAP)
1.4. | Title 40 Code of Federal Regulations Part 3 (40 CFR Part 3) -- the Cross-Media Electronic Reporting Rule (CROMERR) | Must have
2. Support capacity to handle:
2.1. | Diverse personnel of over 30 persons including chemists, microbiologists, technicians, managers and supervisors, administrative clerks, and IT professionals. | Must have
2.2. | An average of 4500 samples per month | Must have
2.3. | An average of 6000 analyses per month | Must have
2.4. | An average of 20000 analytes per month | Must have
2.5. | The system will support electronic signatures at various locations where applicable (e.g. Approval). | Must have
2.6. | The system will support use of Electronic Lab Notebook (ELN) system | Good to have
3. Support data hierarchy requirements:
3.1. | The system will capture and store data by Budget Unit | Must have
3.2. | The system will capture and store data by Program | Must have
3.3. | The system will capture and store data by Site/Locator (Sample Point) | Must have
3.4. | The system will capture and store data by Sample Type | Must have
3.5. | The system will capture and store data by Test | Must have
3.6. | The system will capture and store data by Analyte | Must have
3.7. | The system will capture and store data by Permitted/Not Permitted | Must have
3.8. | The system will capture and store data by sampling event (date/time) | Must have
3.9. | The system will allow for multiple samples to be analyzed for each program/project | Must have
3.10. | The system will allow for samples to be analyzed for multiple analyses | Must have
3.11. | The system will allow for multiple analytes to be reported for each analysis | Must have
4. Configure system-wide data format attributes:
4.1. | The system will store date as YYYY/MM/DD but system can display on users preference (i.e. MM/DD/YYYY) | Must have
4.2. | The system will store and format time in military time as HH:MM | Must have
4.3. | The system will be flexible to recognize and convert non-military time into military time. | Good to have
4.4. | The system will support the creation of user define rounding rules that specify both significant figures and decimal places to be reported (per client, permit, program, concentration, analysis, analyte, or QC type) | Must have
Client & Contact Information
5. Record Client information: |  |  | Client Information
5.1. | Client ID (unique) | Must have
5.2. | Client Description | Must have
5.3. | Client Mailing Address | Must have
5.4. | Client Reporting Address | Must have
5.5. | Client Billing Address | Must have
5.6. | Client Phone Number | Must have
5.7. | Client Fax Number | Must have
5.8. | Client Cell Phone Number | Must have
5.9. | Client Email | Must have
5.10. | Client Job No. and comment (multiple job numbers) | Must have
5.11. | Client messaging details | Must have
5.12. | Client Project Manager details | Must have
6. Record Contact information:
6.1. | Contact Name | Must have
6.2. | Contact Mailing Address | Must have
6.3. | Client Reporting Address | Must have
6.4. | Client Billing Address | Must have
6.5. | Contact Phone Number | Must have
6.6. | Contact Fax Number | Must have
6.7. | Contact Cell Phone Number | Must have
6.8. | Contact Position/Title | Must have
6.9. | Contact Email | Must have
6.10. | Contact messaging details | Must have
6.11. | Contact Project Manager details | Must have
6.12. | The system will allow multiple contacts to be associated with a Client | Must have
7. Client Management functionality:
7.1. | The system will be able to track and capture Client complaints | Must have
7.2. | The system will have the ability to setup automatic notifications to Clients | Must have
7.3. | The system will have a web portal for Clients to access their information (e.g. update their own contact information, access results, etc.) | Nice to have
7.4. | The system will provide a means for Client to perform self-service queries | Must have
7.5. | The system will allow changes to the reports to be made by the end-user | Good to have
7.5.1. | The system will allow the reports to be expanded/updated by the EBMUD staff | Must have
7.5.2. | The system will allow changes to how data is displayed in reports (e.g. unit conversions) | Good to have
7.5.3. | The system will report to specified significant figures and/or decimal places as specified by the regulatory agency, and/or by the method | Must have
7.5.4. | The system will be able to associate cost with Test, Analyte, or group of tests | Must have
7.5.5. | The system will allow adjustment to cost directly or with multipliers | Must have
7.5.6. | The system will be able to associate cost for billing to client by report ID, sampling event or frame of time. | Must have
7.5.7. | The system will be able to generate a quote and Work Authorization for client approval of new projects | Must have
7.5.8. | The system will be able to generate a billing log for invoicing clients. The log will lock costs associated with numbered report. | Must have
7.5.9. | The system will be able to generate a cost forecast reports for clients of their projects for frame of time (calendar year, fiscal year, etc.). | Must have
7.5.10. | The system will be able to generate the cost value of analyses for custom reports to include/exclude in-house projects, clients, time frame, etc.) | Must have
7.5.11. | The system will be able to capture cost for sample collection labor, analytical labor, and reporting labor. | Nice to have
7.5.12. | The system will be able to capture shipping costs, special supplies, and/or equipment needed for project and client. | Must have
7.5.13. | The system will allow user defined parameters for the use of "<", ">", or ">=" values in calculations (i.e., treat "<" value as on-half the value, as zero, absolute value, etc.) | Must have
7.6. | The system will restrict Client-level queries to those projects/programs assigned to the Client | Must have
7.7. | The system will allow external access via a secure portal for specific client data | Nice to have
Program Information & Analysis Templates
8. Record Program information: |  |  | Program Information
8.1. | The system will record the associated Client | Must have
8.2. | The system will record the job number | Must have
8.3. | The system will record the program name | Must have
8.4. | The system will record the program type | Must have
8.5. | The system will record Start Date | Must have
8.6. | The system will record End Date | Must have
8.7. | The system will calculate program duration based on start and end dates | Good to have
8.8. | The system will record the project manager | Must have
8.9. | The system will record the exact sampling locations | Must have
8.10. | The system will record the sampling type (e.g. grab/composite) | Must have
8.11. | The system will be able to have 30+ sampling types | Must have
8.12. | The sampling types will be able to be defined by the user | Must have
8.13. | The system will record sampling frequency by program and parameter | Must have
8.14. | The system will record the methods included in the project | Must have
8.15. | The system will record the analytes measured in the project | Must have
8.16. | The system will record project-specific analytical limits, rounding, reporting | Must have
8.17. | The system will allow special QA requirements per program/permit | Must have
8.18. | The system will record at the project level the bottle type and preservation to be used with samples | Must have
8.19. | The system will record multiple types of final report required (Federal and State forms) | Must have
8.20. | The system will record multiple recipients of the final report | Must have
8.21. | The system will allow comments to be entered against a project | Must have
8.22. | The system will have the ability to link a project to an outside document | Good to have
8.23. | The system will have the ability to schedule receipt of samples pending against a project | Must have
8.24. | The system will have the capability of building a recurring sample schedule for on-going or routine sampling and analysis | Must have
8.24.1. | The system will allow the recurring schedule to repeat indefinitely without user intervention | Must have
8.24.2. | The system will have the capability of removing unreceived samples from the sample schedule, when configured to do so | Must have
8.24.3. | The system will have the ability to change the expect date for the unreceived samples based on Client request/operational changes | Must have
8.24.4. | The system will have the ability to globally change all sample information fields (such as container type, container quantities, analysis method, location) for all scheduled but not yet received samples and program information (i.e. container type for alkalinity analysis changes and all scheduled samples need to be updated with the new container type). | Must have
8.25. | The system will have the ability to create project and program templates that facilitate in creation of common project configurations | Must have
8.26. | The system will link to program maps (editable) | Must have
8.27. | The system will have ability to attach documents and hyperlinks (editable) | Must have
8.28. | The system will have Program corrective action capabilities | Must have
8.29. | The system will have Program reversioning capabilities | Must have
9. Record Analysis Template information:
9.1. | The system will record user-defined templates for analysis (i.e. test codes) | Must have
9.2. | The system will record a name for the analysis code | Must have
9.3. | The system will record a method for the analysis code | Must have
9.4. | The system will record the matrix for the analysis | Must have
9.4.1. | The system will allow multiple test codes to point to a single method | Must have
9.5. | The system will record results for one or more parameters for the test code | Must have
9.5.1. | The system will allow parameters to be indicated as reportable or non-reportable | Must have
9.5.2. | The system will allow calculations to be defined based on parameters in multiple test codes, based on parameter results (e.g. Bacti, hardness, multiple containers) | Must have
9.5.3. | The system will maintain, assign and meet certification requirements for each instrument & piece of equipment | Must have
9.5.4. | The system will maintain a list of valid instruments, analysts (IDOC/DOC) for each test | Must have
9.6. | The system will identify sample and analysis for regulatory reporting | Must have
9.7. | The system will associate the container, based on the analysis, with the location of storage/analysis (i.e. in-house, specific sub-lab, satellite lab) | Must have
9.8. | The system will associate a test with a specific type of container | Must have
9.9. | The system will have the ability to link an analysis/method to an outside document | Must have
9.10. | The system will have the ability to link an analyte to an outside document | Good to have
9.11. | The system will have the ability to create analysis templates that facilitate the creation of common analytical configurations, including: | Must have
9.11.1. | Analysis Name | Must have
9.11.2. | Analysis Section | Must have
9.11.3. | Analytes | Must have
9.11.4. | Analyte STORET numbers | Must have
9.11.5. | CAS No. | Must have
9.11.6. | Method | Must have
9.11.7. | Holding Time/Preservation | Must have
9.11.8. | Certification/Not (Multiple) | Must have
9.11.9. | Units | Must have
9.11.10. | Method Detection Limit (MDL) | Must have
9.11.11. | Detection Limit for Reporting (DLR) | Must have
9.11.12. | Reporting Limit (RL) or Practical Quantitation Limit (PQL) | Must have
9.11.13. | Method Reporting Limit (MRL) | Must have
9.11.14. | Indication of whether result is reportable or non-reportable | Must have
9.11.15. | Calculations associated with analysis | Must have
9.11.16. | Comments (Internal and reportable - editable) | Must have
9.11.17. | Analysis template information required by regulatory agencies | Must have
9.11.18. | Approximate labor hours | Must have
Sampling Location Information
10. Record Sampling Location information: |  |  | Sample Location
10.1. | The system will record the sampling location name | Must have
10.2. | The system will record the description of the sampling location | Must have
10.3. | The system will record the address of the sampling location | Must have
10.4. | The system will allow multiple permit numbers to be associated with a sampling location | Good to have
10.5. | The system will record other identifiers associated with the sampling location: | Must have
10.5.1. | Latitude, Longitude | Must have
10.5.2. | Structure ID | Must have
10.5.3. | Locator name | Must have
10.5.4. | Site Name | Must have
10.5.5. | Surface Water Classification | Must have
10.5.6. | Monitoring Type | Must have
10.5.7. | District Customer ID | Good to have
10.5.8. | Matrix | Must have
10.5.9. | Pressure Zone | Must have
10.5.10. | The system will be able to add other identifiers associated with the sampling location | Must have
10.6. | The system will store the sampling location information for re-use during future sampling events | Must have
Sample Information
11. Capture Sample information: |  |  | Sample Info
11.1. | The system will create a unique sample number identifier where the numbers increase sequentially. | Must have
11.2. | The system will be able to handle 99 aliquots per container with a unique ID and trackable to the parent | Must have
11.3. | The system will be able to handle 999 samples at a single location | Must have
11.4. | The system will record Sample Number and Sample Name | Must have
11.5. | The system will record collection type (e.g. composite, grab, other as defined by user) | Must have
11.6. | Composite (16/24 hr manual or auto sampler) | Must have
11.6.1. | The system will be capable of tracking individual grab samples used for composites | Must have
11.6.2. | The system will record whether the sample is flow-weight composited or collected over pre-set time intervals | Must have
11.6.3. | The system will be able to build Flow Compositing Worksheets where flow by hour is required. | Must have
11.6.4. | The system will record the duration (in hours or days) of the composite sampling | Must have
11.6.5. | The system will record the temperature of the composite samples | Must have
11.7. | Grab | Must have
11.7.1. | The system will record the type of collection | Must have
11.7.2. | The system will record the collection method name | Must have
11.7.3. | The system will record the Equipment (SSB, Peristaltic Pump, Plastic Bucket, etc.) used to collect the grab sample | Must have
11.8. | The system will record sample date range and time | Must have
11.9. | The system will record sample matrix | Must have
11.10. | The system will record sample description (1000 characters) | Must have
11.11. | The system will record sample container information | Must have
11.11.1. | The system will be able to have multiples of the same container for a single site (sample) | Must have
11.11.2. | The system will support multiple containers per sample | Must have
11.11.3. | The system will handle spares that are collected but may not be analyzed | Must have
11.11.4. | The system will record number of containers | Must have
11.11.5. | The system will record sample container type | Must have
11.11.6. | The system will record sample container lot number | Must have
11.11.7. | The system will record sample amount | Must have
11.11.8. | The system will support volume units | Must have
11.11.9. | The system will support weight units | Must have
11.12. | The system will record the sample analysis that was requested | Must have
11.13. | The system will allow for multiple analyses per sample | Must have
11.14. | The system will allow for multiple projects per sample | Must have
11.15. | The system will allow the user to specify a priority (turnaround time (TAT)) | Must have
11.16. | The system will allow the adjustment of prices as TAT is adjusted using a factor if possible | Must have
11.17. | The system will allow the user to add/delete analyses based on role/privileges | Must have
11.18. | The system will allow the user to add /delete samples at any time during the project | Must have
11.19. | The system will require the user to provide an explanation for deletion/cancelation of samples | Must have
11.20. | The system will be allowed to cancel all containers at a single location | Must have
11.21. | The system will record comments at the sample level | Must have
11.22. | The system will record comments at the project level | Must have
11.23. | The system will record comments at the analyte level | Must have
11.24. | The system will have the ability to create user-defined fields and store data in those fields at the sample level | Must have
Chain-of-Custody, Container Labels, & Sample Kits
12. Generate a Chain-of-Custody (CoC): |  |  | COC
12.1. | The system will include an all-electronic/paperless CoC process option | Must have
12.2. | The system will generate a chain-of-custody identifier automatically | Must have
12.3. | The system will provide the sample identifier for each sample on the CoC | Must have
12.4. | The system will allow the user to break up a sampling event into multiple CoCs | Must have
12.5. | The system will provide the aliquot identifier for aliquots included in the CoC | Must have
12.6. | The system will provide the program name/code, if appropriate | Must have
12.7. | The system will provide the sampling information: | Must have
12.7.1. | Sample Site | Must have
12.7.2. | Sample Location | Must have
12.7.3. | Sample Location description for non-routine/unknown samples | Must have
12.8. | The system will provide the sampling agency, if required | Good to have
12.8.1. | The system will provide the address (ID No.) | Good to have
12.8.2. | The system will provide the phone number | Good to have
12.8.3. | The system will provide the fax number | Good to have
12.8.4. | The system will provide the contact person | Good to have
12.8.5. | The system will provide the email address | Good to have
12.9. | The system will provide the sampling program | Must have
12.10. | The system will provide the sampling event date range | Must have
12.11. | The system will provide the sampling event time | Must have
12.12. | The system will provide the sample location | Must have
12.13. | The system will provide the sampling address | Good to have
12.14. | The system will provide the requested analysis per container | Must have
12.15. | The system will provide the sample information | Must have
12.15.1. | Expected sample matrix | Must have
12.15.2. | Grab/Composite | Must have
12.15.3. | Container type with lot number | Must have
12.15.4. | Container size | Must have
12.15.5. | Preservative needed | Must have
12.16. | The system will associate COC, program, kits, and containers | Must have
12.17. | The system will record the number of containers covered by the CoC | Must have
12.18. | The system will record additional comments (report and internal fields) | Must have
12.19. | The system will generate a chain-of-custody that is compliant with NELAC standards. | Must have
13. Generate Container Labels:
13.1. | The system will create a unique container identifier, based on the sample identifier | Must have
13.2. | The system will generate container labels with the following information: | Must have
13.2.1. | The label will display the Container ID | Must have
13.2.2. | The label will display a barcode representation of the sample #/Container ID | Must have
13.2.3. | The label will display the sample location | Must have
13.2.4. | The label will include sample composition information (if applicable) | Must have
13.2.5. | The label will display the requested analysis | Must have
13.2.6. | The label will display the sample matrix | Must have
13.2.7. | The label will display the preservative (If Applicable) | Must have
13.2.8. | The label will contain the bottle type | Must have
13.2.9. | The label will display the collection date and time | Must have
13.2.10. | The label format will be compliant with NELAC standards | Must have
13.2.11. | The label format should contain Recollect Identification (If Applicable) | Must have
13.2.12. | The label format may contain Route, program, permitted Y/N, and client name | Must have
14. Manage Sample Kits:
14.1. | The system will allow for safety alerts for sampling potentially hazardous or industrial sites to notify field sampling personnel and analysts | Must have
14.2. | The system will capture the Kit ID used for sample transport | Must have
14.3. | The system will assign status to Kit ID so that the kit can be tracked during the following phases: scheduled, packed, reviewed, ready for deployment, deployed, received | Must have
14.4. | The system will generate sampling instruction to sampler based on analysis requested in sampling kit | Must have
14.5. | The system will be able track comments associated with Sample Kits | Must have
Sample Collection
15. Capture Sample Collection information: |  |  | Sample Collection
15.1. | The system will record the sampler (name) | Must have
15.1.1. | The system will validate against EBMUD employee directory to populate/auto-complete employee name or employee ID number interchangeably, but also allow for entering non-District employee names | Good to have
15.2. | The system will record that the pH of the samples was checked by the sampler | Good to have
15.3. | The system will record that the chlorine of the samples was checked by the sampler | Good to have
15.4. | The system will record that the conductivity of the samples was checked by the sampler | Good to have
15.5. | The system will record that the water temperature of the samples was checked by the sampler | Good to have
15.6. | The system will record that some other definable field measurement of the samples was checked by the sampler | Nice to have
16. Have In-the-Field mobile app:
16.1. | The mobile application will record the pH checked by the sampler | Nice to have
16.2. | The mobile application will record the chlorine checked by the sampler | Nice to have
16.3. | The mobile application will record the conductivity checked by the sampler | Nice to have
16.4. | The mobile application will record that the water temperature checked by the sampler | Nice to have
16.5. | The mobile application will record other field measurements checked by the sampler | Nice to have
16.6. | The mobile application will record comments and other field notes by the sampler | Nice to have
16.7. | The system will allow the user to indicate one or more malfunction codes for field instruments or meters | Nice to have
Sample Login Information
17. Capture Sample Login information: |  |  | Sample Login
17.1. | The system will record comments at the container level | Must have
17.2. | The system will record Chain-of-Custody number | Must have
17.3. | The system will record the project identifier | Must have
17.4. | The system will record the sample analysis at the time of receipt | Must have
17.4.1. | The system will flag analyses being assigned to containers lacking the required bottle type and/or preservative | Must have
17.4.2. | The system will allow a user to reject a kit or individual samples (e.g. for wrong container, temperature anomaly, cracked bottle, air bubbles) prior to login. | Must have
17.4.3. | The system will allow qualifiers for kit anomalies to be captured during sample receiving | Must have
17.5. | The system will record the sample matrix (water, solid, etc.) | Must have
17.6. | The system will record the pH upon receipt, if required. | Must have
17.7. | The system will record the temperature upon receipt | Must have
17.8. | The system will record sample preservatives per method requirements | Must have
17.9. | The system will record lot# and time of preservation (e.g. metals) | Must have
17.10. | The system will record who preserved the sample (Analyst ID or preserver's name) | Must have
17.11. | The system will record date sample taken | Must have
17.12. | The system will record sample collection time | Must have
17.13. | The system will record date sample received | Must have
17.14. | The system will record time sample received | Must have
17.15. | The system will record who received the samples | Must have
17.16. | The system will allow the user to document that all samples are present | Must have
17.17. | The system will allow the user to indicate sample/container conditions with receiving qualifiers and comments | Must have
17.17.1. | Received frozen | Must have
17.17.2. | Sample container damaged in shipment, no analysis performed | Must have
17.17.3. | Sample container damaged, analysis performed | Must have
17.17.4. | Sample temperature (Targeted sample container in received samples) | Must have
17.17.5. | Thermometer ID | Must have
17.18. | The system will allow each sample/container condition to trigger an action in the system when applied to a sample/container: | Must have
17.18.1. | Continue with analysis | Must have
17.18.2. | Cancel analysis and why | Must have
17.18.3. | Recollect Notes or Comments | Must have
17.19. | The system will allow the user to indicate one or more codes for no analysis | Must have
17.19.1. | No analysis required | Must have
17.19.2. | No analysis due to holiday schedule | Must have
17.20. | The system will allow the user to indicate one or more codes when no sample was taken | Must have
17.20.1. | Low flow or no flow | Must have
17.20.2. | No sample required at this time | Must have
17.20.3. | Scheduled sample not received | Must have
17.20.4. | No sample due to equipment failure | Must have
17.21. | The system will record receiving codes at the container level | Must have
17.22. | The system will record at the container level receipt conditions for containers for analyses that have not met method criteria for sample receipt (i.e. air bubbles in VOA containers) and allow user to assign a qualifier that may appear on the final report with the results | Must have
17.23. | The system will be capable of receiving samples by entering data via keyboard | Must have
17.24. | The system will be capable of receiving samples individually | Must have
17.25. | The system will be capable of receiving samples by CoC | Must have
17.26. | The system will allow a sample received at the primary laboratory to be transferred to another laboratory for testing of all or a subset of parameters | Must have
17.27. | The system will allow receiving codes to be entered against individual or multiple samples containers at the time of receipt | Must have
17.28. | The system will allow comments to be entered against a sample container at the time of sample receipt | Must have
17.29. | The system will record coolant type and condition and if they met method requirements | Must have
Subcontracted Laboratories
18. Manage interaction with Subcontracted Laboratories: |  |  | Subcontract
18.1. | The system will record information about subcontracting laboratories | Must have
18.1.1. | The system will allow to attach files to subcontracted laboratory details | Must have
18.1.2. | The system will record laboratory name | Must have
18.1.3. | The system will flag data from subcontract laboratories that are not NELAC compliant | Must have
18.1.4. | The system will flag data from subcontract laboratories that are client/permit/matrix approved | Must have
18.1.5. | The system will record multiple laboratory contact | Must have
18.1.6. | The system will capture the EBMUD vendor code from purchasing to link to our system to make tracking of costs easier | Must have
18.1.7. | The system will record multiple laboratory address | Must have
18.1.8. | The system will record multiple laboratory phone/fax numbers | Must have
18.1.9. | The system will record multiple laboratory email address | Must have
18.1.10. | The system will record multiple laboratory certification number | Must have
18.2. | The system will record the date sample is sent | Must have
18.3. | The system will record the date sample was received by a subcontract lab | Must have
18.4. | The system will record the analysis requested by method | Must have
18.5. | The system will record the analysis results | Must have
18.6. | The system will record the number of bottles sent | Must have
18.7. | The system will record the sample description | Must have
18.8. | The system will record the sample holding time | Must have
18.9. | The system will record the sample container's preservation method | Must have
18.10. | The system will record the sample preparation date and time relating to the subcontracted analysis | Must have
18.11. | The system will record or calculate the turnaround time for each subcontracted analysis | Must have
18.12. | The system will apply limits to samples analyzed by a subcontractor in the same manner they are applied to samples analyzed within the lab | Must have
18.13. | The system will record the price of the analysis | Must have
18.14. | The system will have the capability to import subcontracted analyses results provided in a variety of electronic formats (e.g. XML, CSV). | Must have
18.15. | The system will support multiple import scripts that can be reused | Must have
18.16. | The system will generate sub-contract CoC using laboratory unique identifiers for the sample | Must have
18.17. | The system will support association of import scripts with a program or subcontract entity | Must have
18.18. | The system will be able to build invoices for subcontract laboratories that can be passed to accounting | Must have
19. Capture Subcontracted Laboratory sample information:
19.1. | The system will record the subcontracted laboratory (Name & ID) that is assigned to perform the analysis | Must have
19.2. | The system will capture the name of the subcontracted test | Must have
19.3. | The system will capture the subcontracted parameter | Must have
19.4. | The system will capture the subcontracted parameter MDL | Must have
19.5. | The system will capture the subcontracted parameter RL | Must have | Subcontract
Sample Management
20. Have Sample Management and Tracking functionality: |  |  | Sample Mgmt
20.1. | The system will track the location of each sample container (i.e. in-house, subcontract lab) | Must have
20.2. | The system will allow for multiple analyses using the same container | Must have
20.3. | The system will allow for multiple containers for a single analysis | Must have
20.4. | The system will record the sample status | Must have
20.4.1. | Scheduled | Must have
20.4.2. | Pre-logged | Must have
20.4.3. | Received | Must have
20.4.4. | Pending Prep | Must have
20.4.5. | In Progress Prep (associated batch) | Must have
20.4.6. | Pending Analysis | Must have
20.4.7. | Completed | Must have
20.4.8. | In Progress Analytical (associated batch) | Must have
20.4.9. | Cancelled | Must have
20.4.10. | Re-sampled | Must have
20.4.11. | Re-analysis | Must have
20.4.12. | Sent (Commercial Lab) | Must have
20.4.13. | Reviewed | Must have
20.4.14. | Final Reviewed | Must have
20.4.15. | Senior Chemist Reviewed | Must have
20.4.16. | Manager Approved | Must have
20.4.17. | Un-reported | Must have
20.4.18. | Re-Reported | Must have
20.4.19. | Disposed | Must have
20.5. | The system will capture a time stamp at each status change to allow duration calculations | Must have
20.6. | The system will have the ability to change the sample status by login, batch, analytical group | Must have
20.7. | The system will record whether a sample is identified as hazardous | Must have
20.8. | The system will record the date the sample is disposed | Must have
20.8.1. | The system will capture who disposed of the sample | Must have
20.8.2. | The system will need to generate a Disposal Status Summary report requested by clients | Good to have
20.9. | The system will indicate the sample's disposal method | Must have
20.9.1. | Un-approved samples will be marked to be exclude from disposal | Must have
20.10. | The system will be able to dispose of sample on the container, login, batch level or user defined requirements | Must have
20.11. | The system will generate a Daily Sample Status report | Must have
20.11.1. | The system will provide the ability to generate the report by any combination of laboratory group/unit and method | Must have
20.11.2. | The report will be sorted and/or filtered by assigned analyst | Must have
20.11.3. | The report will be sorted and/or filtered by lab unit then method and/or analyte | Must have
20.11.4. | The report will include sample date and time | Must have
20.11.5. | The report will include sample ID | Must have
20.11.6. | The report will include sample name | Must have
20.11.7. | The report will include sample location | Must have
20.11.8. | The report will indicate holding time remaining by analyte | Must have
20.11.9. | The report will indicate the holding time by analyte | Must have
20.11.10. | The system will store and display the client due date & associated TAT | Must have
20.12. | The system will allow the user to search and retrieve sample information (ad-hoc reporting) | Must have
20.12.1. | Retrieve sample information based on comments, collection, flow, matrix, analyte, permitted/not permitted, sample collection attributes, reported result value, flag, sample site attributes | Must have
20.12.2. | Retrieve sample information based on Client | Must have
20.12.3. | Retrieve sample information based on program | Must have
20.12.4. | Retrieve sample information based on sampling location | Must have
20.12.5. | Retrieve sample information based on sample type | Must have
20.12.6. | Retrieve sample information based on status | Must have
20.12.7. | Retrieve sample information based on analysis type | Must have
20.12.8. | Retrieve sample information for “overdue” samples | Must have
20.12.9. | Retrieve based on CoC# or sample number | Must have
20.12.10. | Retrieve based on Client sample identifier | Must have
20.12.11. | Retrieve based on Analyst | Must have
20.12.12. | The system will support the tracking of groups of Client samples, such as by worklist. | Must have
20.13. | The system will generate a sample receipt report daily showing submitted samples (Custom report) | Must have
21. Track Aliquots for Sub-sampling:
21.1. | The system will append when needed, an aliquot number to the sample identifier | Must have
21.2. | The system will capture date aliquot was taken | Must have
21.3. | The system will capture time aliquot was taken | Must have
21.4. | The system will capture the person who performed the aliquot | Must have
21.5. | The system will capture what analytes the aliquot is being tested for | Must have
21.6. | The system will maintain Parent/Child relationship | Must have
Scheduling
22. Feature scheduling tools: |  |  | Schedule
22.1. | The system will have comprehensive scheduling tools for all functional aspects of the system | Must have
22.2. | The system will support interval scheduling (once a day, week, month, etc.) | Must have
22.3. | The system will support calendar/absolute date scheduling | Must have
22.4. | The system will have flexible recurring scheduling options, and take into account weekends and holidays | Must have
22.5. | The system will have a calendar view | Must have
22.6. | The system will generate a report showing future schedule of analyses by project/program, client, and analysts for a specified time, test, parameter, and/or analysts | Must have
Sample Preparation
23. Capture Sample Preparation information: |  |  | Sample Prep
23.1. | The system will record preparation method | Must have
23.1.1. | The system will allow for and link multistep preparation methods | Must have
23.2. | The system will have a QC template for the prep batch based on matrix and analytical method | Must have
23.3. | The system will record the analyst who performed the preparation. | Must have
23.4. | The system will record analyst(s) who initiated sample preparation | Must have
23.5. | The system will record date/time and analyst who performed filtration | Must have
23.6. | The system will record/allow for one analyst to start prep and another to finish | Must have
23.7. | The system will record analyst(s) who completed sample preparation | Must have
23.8. | The system will record the date and time started | Must have
23.9. | The system will record the date and time completed | Must have
23.10. | The system will record the starting/ending sample amount | Must have
23.11. | The system will record standards and reagents used during the preparation | Must have
23.12. | The system will record sample multiple dilutions | Must have
23.13. | The system will record comments about the sample preparation process | Must have
23.14. | The system will record all supporting instrumentation (balance, incubator, pipettes, etc.) and allow for multiple instruments | Must have
23.15. | The system will be able to capture incubator temperatures or interface with thermocouples | Good to have
23.16. | The system will generate a unique prep batch number | Must have
23.17. | The system will be able to record prep type (chloro freezing) | Good to have
23.18. | The system will record preparation criteria (e.g. digestion temp, time, pressure) | Must have
Sample Analysis
24. Generate Sample Batches: |  |  | Analysis
24.1. | The system will allow the user to generate batches by method and/or analyte | Must have
24.1.1. | The system will allow user to generate batches with multiple methods (test codes) and/or analyte, e.g. MTF batches have SM 9221B, SM9221E and SM9221F | Must have
24.2. | The system will allow the user to generate batches with multiple sets of QC samples | Must have
24.3. | The system will have a user defined QC template for the analytical batch based on matrix and analytical method | Must have
24.4. | The system will capture the analyst ID that created any batch within the system. | Must have
24.5. | The system will allow an analytical batch to be associated with one or more corresponding prep batches. | Must have
24.6. | The system will allow the user to specify the order of samples within a batch. | Must have
24.7. | The system will allow samples to be selected for batching from a list of pending analyses/preparations | Must have
24.8. | The system will record multiple dilutions | Must have
24.9. | The system will allow the user to sort the list of pending analyses/preparations | Must have
24.9.1. | Sort by holding time | Must have
24.9.2. | Sort by due date | Must have
24.9.3. | Sort by sample type and location | Must have
24.9.4. | Sort by permitted/ not permitted | Must have
24.10. | The system will allow users to add/remove analytes from analysis. | Must have
24.11. | The system will have a visual warning of samples in the list of pending analyses/preparations that are in jeopardy of exceeding their holding time (e.g. color scale) | Must have
24.12. | The system will allow samples to be added or removed from the batch after initial creation | Must have
24.13. | The system will generate a unique batch identifier | Must have
24.14. | The batch will contain the Sample ID | Must have
24.15. | Each QC sample created will have a unique identifier | Must have
24.16. | The batch will contain the Sample Container ID | Must have
24.17. | The batch will contain sample weight | Must have
24.18. | The batch will contain final volume of sample | Must have
24.19. | The batch will contain filtered volume of sample | Must have
24.20. | The batch will contain the spike information | Must have
24.20.1. | Sample identifier of sample spiked | Must have
24.20.2. | Amount spiked | Must have
24.21. | The batch will record comments from the user | Must have
24.22. | The system will print out a benchsheet | Must have
24.23. | The system will print out a batch summary | Must have
24.24. | The system will allow for each QC sample to be referred back to the specific batch | Must have
24.24.1. | The system will allow the user to configure QC samples and define intervals | Must have
24.24.2. | The system will allow the definition of batch templates by analysis, which indicates how QC samples and production samples are to be structured in the batch. | Must have
24.24.3. | The system will allow the user to define QC sample calculations based on user roles and privileges | Must have
24.25. | The system will allow the user to develop an instrument sample sequence from the batch | Nice to have
24.25.1. | The system will allow a sequence created in LIMS to be ported into the instrument software if bi-directional integration is possible | Nice to have
24.25.2. | The sequence will contain the sample number identifier | Nice to have
24.25.3. | The sequence will contain the date of analysis | Nice to have
24.25.4. | The sequence will contain the instrument calibration standards | Nice to have
24.25.5. | The sequence will contain sample type and QC type | Nice to have
24.25.6. | The sequence will contain calibration check standards | Nice to have
24.25.7. | The system will record all supporting instrumentation (balance, incubator, pipettes, etc.) and allow for multiple instruments | Nice to have
24.26. | The system will allow multiple products or a wildcard function when selecting analyses for a batch | Must have
25. Capture Sample Analysis information:
25.1. | The system will capture analysis date | Must have
25.2. | The system will capture analysis time | Must have
25.3. | The system will capture analyst identifier | Must have
25.4. | The system will capture multiple analyst's identifiers for a single analysis | Must have
25.5. | The system will capture the batch ID | Must have
25.6. | The system will capture the process date and time, indicating when results were processed against quality control measures | Must have
25.7. | The system will capture the method name | Must have
25.8. | The system will capture the method reference version (e.g. 18th Edition) | Must have
25.9. | The system will capture the version of the standard operating procedure in effect for the analysis | Must have
25.10. | The system will capture any standards used in the analysis | Must have
25.11. | The system will capture any reagents used in the analysis | Must have
25.12. | The system will allow for analysis of duplicate samples and associate the results of the analysis | Must have
25.13. | The system will record all supporting instrumentation (balance, incubator, pipettes, etc.) | Must have
25.14. | The system will capture calibration data | Must have
25.14.1. | The system will allow calibration data to apply to multiple batches | Must have
25.15. | The system will allow a sample to be reanalyzed | Must have
25.15.1. | The system will store the result of the original analysis and re-analysis | Must have
25.15.2. | The system will retain a reference from multiple re-analysis to the original analysis | Must have
25.15.3. | The system will adjust results based on dilution factor | Must have
25.15.4. | The system will adjust MDL/PQL/RL/MRL limits based on dilution factor | Must have
25.16. | The system will have the capability of qualifying samples that fall outside of user defined acceptance level criteria, as defined by a program | Must have
25.17. | The system will have the capability of notifying that samples fall out of the historical ranges by sample point | Must have
25.17.1. | Calculate the average of a specified date range +/- deviation | Must have
25.18. | The system will record comments related to sample analysis at the batch, method and analyte levels | Must have
25.19. | The system will have the capability of "flagging" results from analyses for which the EBMUD is not accredited. | Must have
25.20. | The system will allow the user to approve data | Must have
25.20.1. | The system will indicate that the data was reviewed/approved by the analyst | Must have
25.20.2. | The system will indicate a final data approval for reporting | Must have
25.20.3. | The system will indicate a Senior Chemist approval for reporting | Must have
25.20.4. | The system will indicate a Laboratory Manager approval for reporting | Must have
25.20.5. | The system will require data to be re-approved, if it is changed after approval. | Must have
25.21. | The system will flag results that do not pass user-specified integrity checks, such as a Total Nitrogen result being less than the Total Kjeldahl Nitrogen result for the same sample | Must have
25.22. | The system will allow for data to become available to our designated users (clients) after final approval | Must have
Data Entry
26. Support Data Entry needs: |  |  | Data Entry
26.1. | The system will support for the entry of mobile, wireless and hardwired (internal) data: | Must have
26.2. | The system will allow data entry in a grid or tabular format (Allow user to order parameters as needed) | Must have
26.3. | The system will allow the user to design tabular data entry forms for each analysis type | Must have
26.4. | The system will allow entry of data from a pre-defined list of values | Must have
26.5. | The system will allow data to be entered manually with spell checking of free text | Must have
26.6. | The system will allow data to be pushed to instruments (e.g. run lists) | Nice to have
26.7. | The system will allow data to be captured from instruments (e.g. balances, meters, probes, GC/MS, BOD analyzer, ELISA plate reader, flow injection analyzer) | Must have
26.8. | The system will support bi-directional communication with ELNs | Nice to have
26.9. | The system will allow the user to define/modify calibration standards and QC samples | Must have
26.10. | The system will allow the user to define calculated fields | Must have
26.11. | The system will support basic calculations | Must have
26.11.1. | The system will have a visual queue for results outside limits or fails | Must have
26.11.2. | The system will support complex calculations with look-up tables capabilities | Must have
26.11.3. | The system will allow the user to define validation rules that confirm impossible values are not entered (e.g. pH = 62, MPN codes greater than 5) | Must have
26.12. | The system will capture calibration data | Must have
26.13. | The system will capture a failure if calibration fails | Must have
26.14. | The system will capture dilution data from both preparation and analysis | Must have
26.15. | The system will allow the user to define data entry templates that can be used for future analysis | Must have
26.16. | The system will allow the user to archive and retrieve the tabular data entry forms | Must have
26.17. | The system will allow supervisors to approve data and add comments | Must have
26.18. | The system will store the final reportable result for all tests, whether entered by the analyst or calculated by the LIMS. | Must have
26.19. | The system will compute results for parameters that are calculated from analytical data that may all be entered from distinct batches | Must have
26.20. | The system will store the raw, un-manipulated/unrounded data as well as the final result | Must have
26.21. | The system will use raw data values (not rounded) in calculations | Must have
26.22. | The system will allow a comment to be added at the parameter, sample, batch, login level | Must have
26.23. | The system will designate if comments are reportable or non-reportable | Must have
26.24. | The system will have the ability to attach photos with report, login, batch | Must have
26.25. | The system will record data entry in real-time, no manual saving required | Must have
26.26. | The system will execute calculations automatically upon result entry | Must have
26.27. | When a result is modified, the system will automatically trigger an update to the calculated value | Must have
27. Capture Organic Extraction data:
27.1. | The system will capture the extraction start/stop date/time | Must have
27.2. | The system will capture the analysis | Must have
27.3. | The system will capture the matrix | Must have
27.4. | The system will indicate if sample is outside of method hold time for extraction | Must have
27.5. | The system will capture the extraction method | Must have
27.6. | The system will capture the extraction equipment | Must have
27.7. | The system will capture the extraction reagents | Must have
27.8. | The system will capture the extraction solvent lot # | Must have
27.9. | The system will capture the surrogate spike solution ID | Must have
27.10. | The system will capture the surrogate concentration | Must have
27.11. | The system will capture the surrogate spike solution volume | Must have
27.12. | The system will capture the internal standard spike solution ID | Must have
27.13. | The system will capture the internal standard concentration | Must have
27.14. | The system will capture the internal standard spike solution volume | Must have
27.15. | The system will capture the analyte spike solution ID | Must have
27.16. | The system will capture the analyte spike concentration | Must have
27.17. | The system will capture the analyte spike volume | Must have
27.18. | The system will capture who spiked the sample | Must have
27.19. | The system will capture the clean-up date | Must have
27.20. | The system will capture the pH | Must have
27.21. | The system will capture the sample volume/weight | Must have
27.22. | The system will capture the final volume | Must have
27.23. | The system will allow for comments | Must have
27.24. | The system will capture “extracted by” (multiple people) | Must have
28. Capture Metals Preparation data:
28.1. | The system will record preparation method | Must have
28.2. | The system will record digestion equipment | Must have
28.3. | The system will record Acids used for digestion & associated lot #s | Must have
28.4. | The system will record who prepared the sample (multiple people) | Must have
28.5. | The system will exclude filtered samples from metals prep batch | Must have
28.6. | The system will record the date/time started | Must have
28.7. | The system will record the date/time completed | Must have
28.8. | The system will record the Start and End temperature | Must have
28.9. | The system will record what the sample was analyzed for | Must have
28.10. | The system will record the sample amount | Must have
28.11. | The system will record pH Check <2 and time performed | Must have
28.12. | The system will record the turbidity of the sample | Must have
28.13. | The system will record spike solution | Must have
28.14. | The system will record final volume | Must have
28.15. | The system will record the digestion vessel | Must have
28.16. | The system will record the digestion vessel before digesting | Must have
28.17. | The system will record the sample weight for solid samples before digesting | Must have
28.18. | The system will record comments | Must have
28.19. | The system will record the initials of person who put sample in storage | Must have
28.20. | The system will record sample storage date | Must have
28.21. | The system will record who reviewed the results | Must have
29. Capture Microbiology data:
29.1. | The system will capture date/time analysis started | Must have
29.2. | The system will capture the temperature of the incubator/waterbaths twice a day, 4 hours apart | Must have
29.3. | The system will capture multiple processing analyst names | Must have
29.4. | The system will capture date/time incubation began | Must have
29.5. | The system will capture date/time incubation ended | Must have
29.6. | The system will capture the raw results (single or multiple counts) | Must have
29.7. | The system will calculate precision criteria from the last 15 samples with positive results of each matrix annually or user specified time interval | Must have
29.8. | The system will determine the Most Probable Numbers from reference tables | Must have
29.9. | The system will capture qualifying comments | Must have
29.10. | The system will report text comments | Must have
29.11. | The system will track number of field samples in a batch and alert the user to number of duplicates required for the batch as per method | Must have
29.12. | The system will evaluate duplicate precision against precision criteria for each analytical batch with duplicate | Must have
29.13. | The system will have the ability to configure qualitative fields (Yes/No) | Must have
29.14. | The system will perform the following calculations | Must have
29.14.1. | Percent Recovery | Must have
29.14.2. | RPD for BOD | Must have
29.14.3. | Average for Heterotrophic Plate Count | Must have
29.14.4. | Percent Weight Loss for Heterotrophic Plate Count | Must have
29.14.5. | The system will calculate Membrane Filter results, CFU/100 mL, from sample volume and raw results | Must have
29.14.6. | The system will calculate MDL based on sample volume | Must have
29.15. | The system will calculate acceptable read date/time range | Must have
29.16. | The system will compare raw values to ideal counting range and add qualifier where required | Must have
29.17. | The system will capture research project results | Must have
29.17.1. | Numerical data | Must have
29.17.2. | Journal entries (comments) | Must have
29.18. | The system will have the ability to qualify QA/QC of prepared media and reagents | Must have
29.18.1. | The system will capture the media/reagent lot # | Must have
29.18.2. | The system will be able to link the media/reagent to the sample ID | Must have
29.19. | The system will have the ability to generate and print labels for prepared media | Must have
29.20. | The system will capture sample dilution data | Must have
29.20.1. | The system will be able to recalculate concentration data based on sample dilution | Must have
29.21. | The system will have the ability to generate QA/QC forms for prepared media | Must have
29.22. | The system will import data from .xls or .txt file type (export from Skalar BOD software) | Must have
29.22.1. | Dilution | Must have
29.22.2. | Sample ID | Must have
29.22.3. | Sample volumes | Must have
29.22.4. | Seed vol. | Must have
29.22.5. | Initial Dissolved Oxygen reading, temperature, date/time | Must have
29.22.6. | Final Dissolved Oxygen reading, temperature, date/time | Must have
29.22.7. | Depletion | Must have
29.22.8. | BOD | Must have
29.22.9. | Average BOD | Must have
29.22.10. | Error codes | Must have
29.22.11. | Seed correction factor | Must have
29.23. | The system will capture raw data and perform calculation of result based on complex sets of logic and mathematics (i.e. BOD, Enterococcus) | Must have
29.24. | The system will integrate with dynamic worksheets for microbiology analyses | Must have
30. Capture Molecular Biology (ELISA) data:
30.1. | The system will capture the sample ID | Good to have
30.2. | The system will capture the assay type | Good to have
30.3. | The system will capture the cell lysis date | Good to have
30.4. | The system will import data from .xls, .txt, or .xml file type (export from SkanIt software) | Good to have
30.4.1. | Well number | Good to have
30.4.2. | Raw absorbance per well | Good to have
30.4.3. | Normalized absorbance per well | Good to have
30.4.4. | Analyte concentration | Good to have
30.4.5. | Standard curve r2 value | Good to have
30.4.6. | CV% (coefficient of variance expressed in %) of replicate absorbance data | Good to have
30.5. | The system will calculate the RPD of duplicate sample | Good to have
30.6. | The system will calculate the % Recovery of spiked sample | Good to have
30.7. | The system will calculate the average concentration of replicate samples | Good to have
30.8. | The system will calculate the CV% of replicate concentrations | Good to have
30.9. | The system will store Minimum Detection Limit value | Good to have
30.10. | The system will store Reporting Limit value | Good to have
30.11. | The system will capture sample storage conditions | Good to have
30.12. | The system will capture sample dilution data | Good to have
30.12.1. | Recalculate captured concentrations based on sample dilution | Good to have
30.13. | The system will capture analyst name | Good to have
31. Capture Bioassay & Toxicology data
31.1. | Test Species | Must have
31.2. | Organism Source | Must have
31.3. | Organism Age | Must have
31.4. | Type of bioassay (chronic or acute; static or renewal) | Must have
31.5. | Method reference | Must have
31.6. | Start Date & Time | Must have
31.7. | End Date & Time | Must have
31.8. | Test concentrations | Must have
31.9. | Most sensitive species | Must have
31.9.1. | Date(s) of last most sensitive species test | Must have
31.9.2. | Date of next most sensitive species test | Must have
31.10. | Statistical Analysis: Test lethal and sublethal endpoints | Must have
31.10.1. | Lowest no effects concentration (NOEC) | Must have
31.10.2. | Point estimate (LC50) and 95% confidence interval | Must have
31.10.3. | Survival of sample | Must have
31.10.4. | Survival of Control | Must have
31.10.5. | Percent Minimum Significant Difference (PMSD) | Must have
31.10.6. | Toxicity units | Must have
31.10.7. | Text Comments | Must have
31.11. | Sample Water Quality: | Must have
31.11.1. | Conductivity (uS/cm or mS/cm) | Must have
31.11.2. | Alkalinity (mg/L) | Must have
31.11.3. | Hardness (mg/L) | Must have
31.11.4. | Salinity (ppt, parts per thousand) | Must have
31.11.5. | Dissolved Oxygen (mg/L) | Must have
31.12. | Daily water quality | Must have
31.12.1. | New/Old only for renewal bioassays | Must have
31.12.2. | Sample concentration | Must have
31.12.3. | pH | Must have
31.12.4. | Temperature (°C) | Must have
31.12.5. | Dissolved Oxygen (mg/L) | Must have
31.12.6. | Salinity (ppt, parts per thousand) | Must have
31.12.7. | Instrument ID | Must have
31.12.8. | Text field | Must have
31.13. | Control Chart: Concentration Coefficient of variation | Must have
31.14. | The system will be able to link to concurrent reference toxicant analyses | Must have
31.15. | The system will have the ability to link multiple samples to one analysis | Must have
31.16. | The system will be able to link multiple analyses with one sample | Must have
31.17. | The system will be able to track reference toxicant stock information associated with analysis | Must have
31.17.1. | Lot Number Format to be determined (administrator defined) | Must have
31.17.2. | Chemicals used along with vendor and lot# | Must have
31.18. | Water Quality Information for reference toxicant test | Must have
· Mortality
· Conductivity (uS/cm or mS/cm)
· Dissolved Oxygen (mg/L)
· pH
· Temperature (˚C)
31.19. | The system will track the food source associated with an analysis | Must have
31.20. | The system will be able to link to the associated field observations | Must have
31.21. | The system will track additional instrument ID associated with each test | Must have
31.22. | The system will capture comments | Must have
Data Review
32. Feature Data Review functionality: |  |  | Data Review
32.1. | The system will be able to allow 4+ stages of review of data (analyst, supervisor, Senior Chemist, Lab Manager) | Must have
32.2. | The system will record the lab staff who reviewed/approved data | Must have
32.3. | The system will allow the user to review data on the screen | Must have
32.4. | The system will allow the user to review data on a printed/electronic report | Must have
32.5. | The system will allow multiple limits to be associated with parameters. | Must have
32.6. | The system will allow the user to view limits associated with the analyses during data review | Must have
32.7. | The system will allow the user to view tests that are scheduled but not complete and their status | Must have
32.8. | The system will allow the user/requestor to mark a sample for reanalysis/recollection and display notification of reanalysis/recollection | Must have
32.9. | The system will allow the user to request an additional analysis | Must have
32.10. | The system will allow the user to view comments from data entry | Must have
32.11. | The system will allow the user to enter additional comments | Must have
32.12. | The system will have the capability of qualifying samples with results in exceedance (i.e. falling outside of user-defined acceptance criteria). | Must have
32.13. | The system will allow the user to indicate the data is approved (final) | Must have
32.14. | The system will allow the user to add a Permitted/Not Permitted flag | Must have
32.15. | The system will display a user defined # of results for samples taken from that sampling location, based on # of points or date range | Must have
33. Support Data Qualifiers and Receiving Qualifiers:
33.1. | The system will include descriptions for all data qualifiers and receiving qualifiers | Must have
33.2. | The system will allow the user to create and define data qualifiers and receiving qualifiers, and set logic for assignment | Must have
34. Review of Quality Control Data:
34.1. | The system will display all associated QC information with a sample number | Must have
34.1.1. | The system will show the batch identifier of the QC samples | Must have
34.1.2. | The system will show the analyst identifier | Must have
34.1.3. | The system will show all related QC samples | Must have
34.1.4. | The system will show the analysis performed on the batch | Must have
34.1.5. | The system will show the parameters | Must have
34.1.6. | The system will show the results | Must have
34.1.7. | The system will show spike recoveries | Must have
34.1.8. | The system will have the ability to flag and notify when an analysis is out of control | Must have
34.1.9. | The system will have the ability to log and track investigations and corrective actions | Must have
34.1.10. | The system will show relative percent difference (RPD) | Must have
34.1.11. | The system will show relative standard deviation (RSD) | Must have
34.1.12. | The system will show critical range for microbiology | Must have
34.1.13. | The system will calculate and report absolute value of difference between logarithmic sample A and sample B (method specific) | Must have
34.1.14. | The system will show average value of replicates | Must have
34.2. | The system will force the user to address QC failures before data review/approval. can continue | Must have
34.3. | The system will allow the ability to evaluate the following quality control parameters | Must have
34.3.1. | Duplicate sample results and their average | Must have
34.3.2. | Percent recovery on QC | Must have
34.3.3. | Percent recovery on MS/MSD | Must have
34.3.4. | Relative Percent Difference on MS/MSD | Must have
34.3.5. | Positive control | Must have
34.3.6. | Negative control | Must have
34.3.7. | BEG blank (for membrane filtration) | Must have
34.3.8. | MID blank (for membrane filtration) | Must have
34.3.9. | END blank (for membrane filtration) | Must have
34.3.10. | Filter blank (for membrane filtration) | Must have
34.3.11. | Blank (for MPN) | Must have
34.3.12. | Dilution factor/initial sample amount | Must have
34.3.13. | Sterility check | Must have
34.3.14. | Statistical analysis of QC data | Must have
· Mean
· Mean % recovery
· Mode
· Geometric mean
· Median
· Standard deviation
· Mean Relative % difference (RPD)
· % Recovery (including surrogates)
34.3.15. | The system will be able to perform statistical calculations of control limits based on historical data. | Must have
34.3.16. | MDL or current acronym | Must have
34.3.17. | ML/RL | Must have
34.3.18. | MRL | Must have
34.3.19. | PQL or current acronym | Must have
34.3.20. | The system will provide the ability to define QC rules. (editable/versionable) | Must have
34.4. | The system will capture the QC data for all approved methods | Must have
34.5. | The system will be able to round (significant figures) based on concentration | Must have
34.6. | The system will allow calculations using log functions | Must have
34.7. | The system will allow the parameter criteria to be set by analysis, by matrix, by sample type (field, QC) | Must have
34.8. | The system will allow MDL calculation by parameter, by # samples, by type of samples | Must have
34.9. | The system will allow reversioning of analysis MDL and MRL. The system will be configurable or, better yet, regularly updated to current regulations on detection and reporting limits calculation (e.g. MUR for wastewater, EPA methods). | Must have
34.10. | The system will have a MDL calculation tool using the current options for MDLs, including the low level fortified standard and method blanks per CFR 136.6 MUR 2017 | Must have
Sample Results & Reporting
35. Feature Sample Results & Reporting functionality: |  |  | Results
35.1. | The system will export the following sample data | Must have
35.1.1. | Sample Identifier | Must have
35.1.2. | Sample Location | Must have
35.1.3. | Sample Site | Must have
35.1.4. | Testing Lab | Must have
35.1.5. | COC number | Must have
35.1.6. | Date Collected | Must have
35.1.7. | Time Collected | Must have
35.1.8. | Date Received | Must have
35.1.9. | Time Received | Must have
35.1.10. | Date of Sample Prep | Must have
35.1.11. | Time of Sample Prep | Must have
35.1.12. | Analyst(s) for Sample Prep | Must have
35.1.13. | Date/Time Analyzed | Must have
35.1.14. | Analyst(s) for Analysis | Must have
35.1.15. | Permit/Program Name | Must have
35.1.16. | Prep/Analytical Batch Number | Must have
35.1.17. | Analytical Method | Must have
35.1.18. | Sample Prep Method | Must have
35.1.19. | STORET code | Must have
35.1.20. | Test/Analyte | Must have
35.1.21. | Detected (“Y” or “N” format) | Must have
35.1.22. | Numeric result (Results like "<2" are stored as "<2" for use in calculations and data exports) | Must have
35.1.23. | Fully Qualified Result (Results like "<2" appear as "<2") | Must have
35.1.24. | Results labeled as "Presence" or "Absence" | Must have
35.1.25. | Result qualifier | Must have
35.1.26. | Accredited parameter ("Y" or "N" format) | Must have
35.1.27. | The system will have the ability to separate the result or lab results qualifier from the result itself (e.g. "<3.2" becomes "<" and "3.2") | Must have
35.1.28. | Result unit | Must have
35.1.29. | Method detection limit (MDL) | Must have
35.1.30. | MDL units | Must have
35.1.31. | Quantitation limit | Must have
35.1.32. | Quantitation limit units | Must have
35.1.33. | PQL limit | Must have
35.1.34. | PQL limit units | Must have
35.1.35. | MRL limit | Must have
35.1.36. | MRL limit units | Must have
35.1.37. | RL limit | Must have
35.1.38. | RL limit units | Must have
35.1.39. | Include results in reports (“Y” or “N” format) | Must have
35.1.40. | Quality Control Type, such as Blank, Calibration Blank, Duplicate | Must have
35.1.41. | Basis (“Wet” or “Dry” basis) | Must have
35.1.42. | Dilution factor | Must have
35.1.43. | Percent recovery | Must have
35.1.44. | RPD | Must have
35.1.45. | Critical Range | Must have
35.1.46. | Sub Sample amount | Must have
35.1.47. | Sub Sample amount unit | Must have
35.1.48. | Initial volume | Must have
35.1.49. | Initial volume unit | Must have
35.1.50. | Final volume | Must have
35.1.51. | Final volume unit | Must have
35.1.52. | Comments | Must have
35.1.53. | Flag indicating data reviewed by QA Group | Must have
35.1.54. | QA Date | Must have
35.1.55. | QA Comment | Must have
35.1.56. | QA Staff User Identifier | Must have
35.1.57. | Kit IDs | Must have
35.1.58. | Receiving Temperatures | Must have
35.1.59. | Preservation | Must have
35.1.60. | Collection Method | Must have
35.1.61. | Testing Lab Certification ID number | Must have
35.1.62. | CAS No. | Must have
35.1.63. | Results ">=" | Must have
35.1.64. | Results below MDL as MDL with U qualifier | Must have
35.1.65. | Client and associated information (e.g. job number) | Must have
35.1.66. | Sample site attributes | Must have
35.1.67. | Program attributes | Must have
35.1.68. | The system will allow the user to define query fields | Must have
35.1.69. | The system will allow the user to generate and save export queries | Must have
35.1.70. | The system will allow users to query by 'Expected Date' for planning. | Must have
35.1.71. | The system will display queries on monitor | Must have
35.2. | The system will have the following reporting capabilities: | Must have
35.2.1. | The system will allow the user to report only project-specific analytes | Must have
35.2.2. | The system will allow reporting of multiple dilutions with corrected MDL/PQL/ML/MRL | Must have
35.2.3. | The system will allow reporting of the wet/dry weight basis for each test | Must have
35.2.4. | The system will indicate which results, if any, in a report are preliminary (i.e. yet to complete the data review process) | Must have
35.2.5. | The system will allow reporting of historical trending for single locations and/or parameters | Must have
35.2.6. | The system will allow reporting of comparative trending between two or more locations and/or analytes | Must have
35.2.7. | The system will allow reporting of statistical analysis of analytical results, including: | Must have
· Standard deviation
· Historical trending
· Geometric mean
· Mean
· Average
· One sided Manns Kendall
· TSI
· WQI
· Half MDL
· At MDL
· At 0
35.2.8. | The system will be able to pull multiple report within the system as a final Client report | Must have
35.3. | The system will indicate on all reports the dates of the original and reissued reports | Must have
35.4. | The system will indicate what information has changed on a reissued report | Must have
36. Generate Electronic Data Deliverable (EDD) compatible with:
36.1. | STORET report and the STORET replacement, WIN | Must have
36.2. | EZ DMR system | Must have
36.3. | DW Report | Must have
36.4. | CIWQS, PET Tool | Must have
36.5. | SMARTS | Must have
Quality Assurance & Control
37. Generate Control Charts: |  |  | QA & QC
37.1. | The system will allow control standards to be charted for each method/analyte | Must have
37.1.1. | The system will accommodate negative controls | Must have
37.1.2. | The system will accommodate positive controls | Must have
37.2. | The system will provide the ability to define the QC data points for evaluation of data | Must have
37.3. | The system will provide the ability to define the QC rules (such as user-specified or statistically derived control limits) | Must have
37.4. | The system will plot the mean line on graph | Must have
37.5. | The system will allow the user to manually enter data | Must have
37.6. | The system will allow for the automatic data entry from instruments | Must have
37.7. | The system will allow for the control limits to be based on plotted data | Must have
37.8. | The system will allow for the control limits to be user defined and displayed on graph | Must have
37.8.1. | The system will allow Lower Control Limits to be defined and displayed on graph | Must have
37.8.2. | The system will allow Upper Control Limits to be defined and displayed on graph | Must have
37.9. | The system will produce charts on a log-log axis | Must have
37.10. | The system will produce charts on a log-linear axis | Must have
37.11. | The system will allow the user to establish alert criteria | Must have
37.12. | The system will produce Individual-X/Moving Range Charts | Must have
37.13. | The system will produce X-Bar and Range Charts | Must have
37.14. | The system will produce Bar Charts | Must have
37.15. | The system will produce Line Charts with mean line | Must have
37.16. | The system will produce Pie Charts | Must have
37.17. | The system will have the ability print all charts and graphs | Must have
37.18. | The system will allow the user to produce charts by various parameters | Must have
37.18.1. | Analyst | Must have
37.18.2. | Instrument | Must have
37.18.3. | Sample type | Must have
37.18.4. | Sample point | Must have
37.18.5. | Matrix | Must have
37.18.6. | Date range | Must have
37.18.7. | Method | Must have
37.18.8. | Procedure | Must have
37.18.9. | Analyte | Must have
37.18.10. | Client | Must have
37.18.11. | Program | Must have
37.19. | The system will have the ability to save queries used to generate control charts for future reuse. | Must have
37.20. | The system will summarize statistics | Must have
37.20.1. | The summary will include recoveries | Must have
37.20.2. | The summary will provide precision | Must have
37.20.3. | The system will be able to calculate and update criteria automatically (with QA review/approval) by test, parameter, method, and matrix with the option to specify a date range (maintain historical results) | Must have
37.20.4. | The system will alert appropriate personnel when limits have been updated. | Must have
37.20.5. | The system will have the ability to flag and notify when an analysis is out of control and will be taken off line until an investigation can be conducted (seven recoveries on one side of the mean, applied to recoveries of QC, etc.) | Must have
37.20.6. | The system will calculate upper and lower control limits(mean +/- 3 std dev) and upper and lower warning control limits (mean +/- 2 std dev) | Must have
37.20.7. | The system will flag data outside +/- 4 std dev for possible removal when calculating limits (huge error rule) | Must have
37.20.8. | The system will allow for ranges to be set by method | Must have
37.20.9. | The system will calculate microbiology critical control and warning limits | Must have
37.21. | The system will be able to have add-on module for complex control charts such as NWA | Must have
37.22. | The X axis of the chart shall display the last X points for sample type, matrix and test combination; a time value of each point can be displayed | Must have
37.23. | the Y axis of the chart shall display the numeric value of a data point being displayed and calculated values from data points as defined for sample type, matrix and data point combination | Must have
38. Track Proficiency Testing (PT):
38.1. | The system will allow scheduling/prelogging and requesting of PT samples by method | Must have
38.2. | The system will record the acceptance ranges associated with the reported results based on the PT provider's evaluations | Must have
38.3. | The system will allow the PT Provider's evaluation to be electronically imported | Must have
38.4. | The system will allow for prelogging or requesting of each PT sample by analyte, method and/or standard operating procedure (SOP) | Must have
38.5. | The system will report the status of all PT's sorted by analyte, method, and/or SOP | Must have
38.6. | The system will generate reports with reported result, analyst, analytical date, sample prep and determinative methods, and analyte(s) | Must have
38.7. | The system will generate reports to include analytes, reported result, and acceptance range | Must have
38.8. | PT Results (% Passed, first or second attempt) | Must have
39. Track Demonstration of Capabilities (DOCs)
39.1. | The system will be able to track by analyst, method, analyte, date range | Must have
39.2. | Separate criteria for initial and continuing | Must have
39.3. | The system will allow to prelog DOC samples (for initial DOC) | Must have
39.4. | The system will notify user defined personnel when training records are nearing expiration date and/or expired | Must have
39.5. | The system will allow to report on DOC (cover sheet, supporting data) | Must have
39.6. | The system will calculate acceptance limits for DOCs and define pass/fail criteria | Must have
39.7. | The system will allow for DOCs by work cell | Must have
39.8. | The system will automatically generate data set for continuing DOCs (e.g. last 4 RE/LRs, LCS, spikes, or other duplicates) by analyst, work cell or method/analyte | Good to have
39.9. | The system will allow for DOC approval at the supervisor and QAO levels | Must have
39.10. | The system will calculate mean and std dev and determine whether bias is high or low | Must have
39.11. | The system will track training per analysis, per analyst, per matrix based on IDOC, PT | Must have
39.12. | The system will be able to update training records by analyst, by analysis, by matrix (allowing multiple matrices to be selected) | Must have
39.13. | The system will be able to manage user training records for verification of current training and certification of users to execute, review and/or approve test procedures | Must have
39.14. | The system will prevent the creation of batches, performance of analysis, and entry of test results if the person has not been trained on the test method or if the training certification has expired | Must have
39.15. | The system will track whether an analyst who is trained on a specific SOP has reviewed newly updated SOP prior to creating a new analytical or preparation batch | Nice to have
39.16. | SOP numbers are linked to test methods in training database | Must have
39.17. | Training record is updated automatically for DOC when a PT sample is analyzed or four consecutive spike blanks were performed and pass QC | Must have
39.18. | The system will have the ability to extract employee overtime hours worked from District ETS and merge with employee overtime hours offered (manually tracked field). The system will display the employees ranked by the addition of these two fields and filtered according to valid training records. | Nice to have
39.19. | The system will allow IDOC/training records for multiple matrices if preparation/analytical steps are identical, e.g., water and wastewater for anions. | Must have
39.20. | The system will issue an automatic warning message to Analyst/lab staff when an SOP is updated in the system | Nice to have
Instrument Management
40. Integrate laboratory instruments: |  |  | Inst. Mgmt
40.1. | The system will allow instruments to be integrated by the EBMUD without involvement of an outside party | Must have
40.2. | The system will allow existing integration with instruments to be modified by the EBMUD without involvement of the LIMS vendor | Must have
40.3. | The system will allow the upload of data into LIMS in the following, but not limited to txt, csv, etc. | Must have
40.4. | The system will allow for creation of a sample sequence with sample preparation information for upload into instrument software | Nice to have
41. Track instrument maintenance:
41.1. | The system will record instrument/equipment Identification (ID) information including manufacturer, serial number, instrument location, date in service, model number, purchase year, service provider, and service provider contact | Good to have
41.2. | The system will record the date the maintenance was performed | Good to have
41.3. | The system will record the person performing the maintenance (internal personal and/or vendor technician) | Good to have
41.4. | The system will record the instrument or equipment identifier on which maintenance was performed | Good to have
41.5. | The system will allow users to select from a pre-defined list of maintenance actions | Good to have
41.5.1. | The system will restrict the pre-defined list of actions to only what is appropriate for the selected instrument | Good to have
41.6. | The system will allow users to enter free-form text describing the maintenance performed | Good to have
41.7. | The system will track when each instrument is in or out of service. | Good to have
41.8. | The system will alert the analyst/coordinator when an instrument has upcoming maintenance (timeframe determined by lab) | Good to have
41.9. | The system will not allow an instrument to be selected by an analyst if under maintenance or overdue for maintenance | Good to have
41.10. | The system will provide notifications for instrument maintenance | Good to have
41.11. | The system will track instrument Service Contract/Agreements | Good to have
42. Capture pipette verification data:
42.1. | The system will capture pipette ID | Good to have
42.2. | The system will capture calibration date | Good to have
42.3. | The system will capture target weight(s) | Good to have
42.4. | The system will capture actual weight(s) | Good to have
42.5. | The system will capture the person performing the calibration | Good to have
42.6. | The system will capture temp of water | Good to have
42.7. | The system will capture the balance ID and weights | Good to have
42.8. | The system will calculate % coefficient of variation, Standard deviation, % relative Standard deviation, bias, acceptance limits | Good to have
43. Capture balance calibration data:
43.1. | The system will capture balance ID | Good to have
43.2. | The system will capture calibration date | Good to have
43.3. | The system will capture target weight(s) | Good to have
43.4. | The system will capture actual weight(s) | Good to have
43.5. | The system will capture the person performing the calibration | Good to have
43.6. | The system will check actual weight against acceptance range | Good to have
43.7. | The system will capture the weight ID | Good to have
44. Capture thermometer calibration data:
44.1. | The system will capture the thermometer id | Good to have
44.2. | The system will capture calibration date | Good to have
44.3. | The system will capture target temperature(s) | Good to have
44.4. | The system will capture actual temperature(s) | Good to have
44.5. | The system will capture the person performing the calibration | Good to have
44.6. | The system will check actual temperature against acceptable range | Good to have
44.7. | The system will also cover NIST thermometer(s) | Good to have
45. Capture Refrigerator, Incubator, Freezer, and Oven (RIFO) data:
45.1. | The system will be able to track daily temperature reading from refrigerators, freezers, ovens and incubators (RIFO) | Good to have
45.2. | The system will push notifications when RIFO has an outlier | Good to have
45.3. | The system will allow the user to edit RIFO thermometer entries | Good to have
45.4. | The system will include worksheet to track and store daily temperature readings from refrigerator, incubators, freezers and ovens (RIFO). | Good to have
45.5. | The system will link the calibration of thermometers (correction factor) with the RIFO tracking | Good to have
46. Contain Instrument Management reports:
46.1. | Report: Maintenance: history, past due, etc. | Good to have
46.2. | Report: Calibration history, past due, etc. | Good to have
46.3. | Report: Instrument usage/utilization by user, method, etc. | Good to have
Consumables Management
47. Feature Consumables Management functionality: |  |  | Consume Mgmt
47.1. | The system will provide a comprehensive set of reports for consumables, summarizing such information like expiration, usage levels, etc. | Must have
47.2. | The system will allow for different status types for consumable status: open, in use, expired, out of stock, etc. | Must have
48. Capture incoming consumable information from vendors:
48.1. | The system will capture be able to file import information from vendor (such as CSV, txt) | Nice to have
48.2. | The system will record the date received | Must have
48.3. | The system will record the number of items and units received | Must have
48.4. | The system will record the standard type | Must have
48.5. | The system will have allow existing reagents/standards to be used to create new types used in the laboratory | Must have
48.6. | The system will record the compounds in the standard | Must have
48.7. | The system will allow a link to MSDS sheets from file server | Must have
48.8. | The system will record the volume/amount | Must have
48.9. | The system will assign a unique identifier | Must have
48.10. | The system will record the vendor | Must have
48.11. | The system will record the vendor’s catalog number | Must have
48.12. | The system will record the vendor’s lot number | Must have
48.13. | The system will record the vendor’s expiration date | Must have
48.14. | The system will record comments | Must have
48.15. | The system will record the date the standard was placed into service | Must have
48.16. | The system will record the date the standard was taken out of service | Must have
48.17. | The system will record the expiration date | Must have
48.18. | The system will record the extension of an expiration date and all associated details (person, date, etc.) | Must have
48.19. | The system will print labels/barcodes with user defined info | Must have
48.20. | The system will be able to record concentration | Must have
48.21. | The system will record the storage location | Must have
48.22. | The system will be able to attach a file (picture, supporting documentation, Certificate of Analysis) | Must have
48.23. | The system will have fields for microbiology genus/species (controls) | Must have
48.24. | The system will indicate storage criteria (frozen, refrigerated, etc.) | Must have
48.25. | The system will communicate with ELN | Good to have
48.26. | The system will record the mode of reagent/material disposition, date/time, and by whom | Must have
48.27. | The system will allow user defined notifications (i.e. expiration date) | Must have
48.28. | The system will record test and section where reagent will be used | Must have
48.29. | The system will have the ability to print labels small enough to fit on a 2 mL vial | Must have
48.30. | The system will record the bacteria strain of control culture stock | Must have
49. Manage Reagent, Media, and Standard Preparations:
49.1. | The system will record the date of preparation | Must have
49.2. | The system will record the starting material | Must have
49.2.1. | The system will record the reference stock identifier | Must have
49.2.2. | The system will record the parent standard used to create intermediate standards | Must have
49.2.3. | The system will record the parent standard used to create working standards | Must have
49.2.4. | The system will record amount used (e.g. g or mL, etc.) | Must have
49.2.5. | The system will record the analytes, if appropriate | Must have
49.2.6. | The system will indicate if the reagent/standard is hazardous | Must have
49.3. | The system will record the stock diluents | Must have
49.3.1. | The system will record the manufacturer | Must have
49.3.2. | The system will record the lot number | Must have
49.4. | The system will record the final volume | Must have
49.5. | The system will record pH | Must have
49.6. | The system will record the concentration of each analyte, if appropriate | Must have
49.7. | The system will record the expiration date | Must have
49.8. | The system will record the extension of an expiration date and all associated details (person, date, etc.) | Must have
49.9. | The system will record the product standard | Must have
49.10. | The system will create a unique reagent/media identifier | Must have
49.11. | The system will record "Prepared by/date" | Must have
49.12. | The system will record "Opened by/date" | Must have
49.13. | The system will record the mode of reagent/material disposition, date/time, and by whom | Must have
49.14. | The system will record spike solution | Must have
49.14.1. | The system will have the capability to record multiple standards used to comprise the spike or spike solution | Must have
49.14.2. | The system will record the components of the spike solution | Must have
49.14.3. | The system will record the spike solution concentration | Must have
49.14.4. | The system will record the spike solution volume used | Must have
49.15. | The system will record comments | Must have
49.16. | The system will have the capability to create a new reagent/standard/media by copying an existing record and altering the pertinent information, and allow error check for duplication | Must have
49.17. | The system will indicate visually when reagents/standards/media are approaching expiration or are expired. | Must have
49.18. | The system will notify the user of upcoming expiration dates | Must have
49.19. | The system will link the reagent/standard/media to appropriate SDS information external to the system. | Must have
49.20. | The system will allow user defined notifications | Must have
49.21. | The system will record test and section where reagent will be used | Must have
49.22. | Reagents, standards, and media can be checked for validity (expiration and qualification status) prior to use | Good to have
49.23. | The system can automatically calculate expiration dates for standards, reagents, and media. | Good to have
49.24. | The system will associate the parent control culture stock used in the creation of daughter culture | Must have
49.24.1. | The system will track each pass a parent control culture stock is used to create daughter culture | Must have
49.24.2. | The system will allow user to define the maximum passes for each parent control culture stock | Must have
49.24.3. | The system will prohibit user to associate the control culture stock to create new daughter culture when the maximum number of passes is reached | Must have
Container Management
50. Feature Container QC functionality: |  |  | Container Mgmt
50.1.1. | The system will allow user defined container types with description | Good to have
50.1.2. | The system will track the following for each container type: vendor name, vendor lot number, C of A availability | Good to have
50.1.3. | The system will track the lot numbers of multiple components that make up a container | Good to have
50.1.4. | The system will track status of container: ordered, received, quarantine, passed QC | Good to have
50.1.5. | For quarantined containers, the system will define and request all user defined QC tests | Good to have
50.1.6. | The system will link supporting documents with container type | Good to have
50.1.7. | The system will notify staff of containers that pass all QC tests and are ready for deployment | Good to have
50.1.8. | The system will recognize future containers with lot numbers which have passed QC and are ready for deployment upon receipt | Good to have
Financial Management
51. Feature financial management tools: |  |  | Fin Mgmt
51.1.1. | The system will have the ability to maintain a Price Schedule | Must have
51.1.2. | The system will have Quotation & Invoice functionality and associated tracking | Must have
51.1.3. | The system will have the ability to generate various financial analysis reports. | Must have
Audit Trail
52. Feature Audit Trail functionality: |  |  | Audit
52.1.1. | The system will record the date the information was changed | Must have
52.1.2. | The system will record the time the information was changed | Must have
52.1.3. | The system will record the user/person that made the changes, requiring proper approvals as needed and record who was the approver | Must have
52.1.4. | The system will retain the original value | Must have
52.1.5. | The system will record the new/adjusted value | Must have
52.1.6. | The system will record the removal or deletion of information, such as canceled analyses or samples | Must have
52.1.7. | The system will require the user to record a reason for the change | Must have
52.1.8. | The system will not allow the audit trail function to be disabled, deleted, or altered | Must have
52.1.9. | The system will allow the system administrator to configure the level of auditing and corresponding "reasons for change" required of users with QAO approval | Must have
52.1.10. | The system will be able to generate an audit trail report (user defined) | Must have
52.1.11. | The system will track changes of approved data and require proper approval (e-signature) | Must have
· Each data change will have a unique ID relating to a form that documents the reason and supporting documents.
· System will prevent the user from using a previously assigned unique ID to make changes to another sample
52.1.12. | The system administrator will be able to configure what level of auditing is required on a table and field basis | Must have
52.1.13. | Any changes made by the system administrator will also be stored in the audit trail | Must have
52.1.14. | The system will ensure that only authorized individuals can create, modify or delete a record | Must have
52.1.15. | The system will ensure that only authorized individuals can create, modify or delete a record | Must have
52.1.16. | The system’s audit trail can be archived | Must have
52.1.17. | All changes on user’s permissions will be recorded in the audit trail | Must have
[7 empty] | Solution Type | Availability | Priority
[7 empty] | Solution NOT Supported | Solution Not supported | Must Have
[7 empty] | Out-of-Box | Future Release | Good to Have
[7 empty] | Add-on-Package | Availabile Now | Nice to Have
Customization